## Slide 1
NEURAL PROSTHETICS BASED ON MICROELECTRODE ARRAYS
Micah Rhea Miriam PINTO

## Slide 2
NEURAL PROSTHETICS BASED ON MICROELECTRODE ARRAYS
MASTER THESIS PRESENTATION – 2024
Pr. Gaëlle LISSORGUES
Pr. Lionel ROUSSEAU
Pr. Matteo COCUZZA
(Dr,HDR)
(Dr,HDR)
Industrial Supervisor
Industrial Co-Supervisor
Academic Supervisor

## Slide 3
Introduction | 01
Research Objectives | 02
Methodology | 03
Micro Fabrication process | 04
Experimental Analysis | 05
COMSOL Simulation | 06
Experiment & Simulation Results | 07
Equivalent Curve Fitting | 08
Longevity Test | 09
Conclusions & Future work | 10
Thesis Presentation Outline

## Slide 4: Introduction
Literature overview
Neural Implants and Prostheses
Overview of Neuroprosthetic Implants.
Overview of Neuronal Communication
Project Neurodiam (2018-2023)
Reviewing prior research and analysing previous studies.
Research objectives and focus
Focus of the thesis with its relevant model.
Key Materials and Techniques
Interdigitated Electrodes.
Electrochemical Impedance Spectroscopy
Equivalent Circuit fit
01
02
03
04

## Slide 5: Overview of Neural prosthetic Implants
Neural Implants and Prostheses
1
2
3
4
Neuro-prosthetic implants are man-made medical devices or simply technical systems that restore sensory or motor functions after injury and modulate neural behavior in neuronal diseases [1].
They are an interface between the body and the machine.
Their purpose is to restore neurological functions by recording or the
stimulation of neurons.
The principal requirements of all medical prosthetic implants are
biocompatibility, bio-adhesion, bio functionality, corrosion resistance,
processability and availability.
An ideal neural implant should minimize the damage to the tissue
and perform reliably and accurately for long periods of time [3].
Figure 1 : Concept of the epileptic seizure detection and suppression using neural recording and stimulation circuits [2].

## Slide 6: Overview of Neuronal Communication
Neural Implants and Prostheses
1
2
3
4
A neuron is an individual nerve cell that makes up the nervous system. It allows information to travel throughout the human body neurons communicate with each other by relaying messages throughout our body powering all our thoughts and actions [4]
Neurons talk to each other using both- electrical and chemical signals.
Messages start as electrical signals travelling rapidly down a neuron.
These signals are called action potentials.
When they reach the gap between two neurons called a synapse, the
information is transformed from an action potential into a chemical
message which crosses the gap. The release of those chemical
messengers can trigger an action potential in the neuron on the other
side of the synapse, conveying the message onward, or it can quiet it.
Each action potential lasts about 1-2 milliseconds. So, a frequency of
1000 Hz is significant because it represents the upper limit of the
theoretical maximum firing rate of neurons, given the typical duration
of action potentials (1 millisecond).
Figure 2 : Anatomy of a neuron [5]
Figure 3 :The action potential graph

## Slide 7: Review of Prior Research
Project Neurodiam (2018-2023)
1
2
The project, which was launched in 2018, aimed to fabricate and optimize a full Diamond implant on a microscale with conducting Boron Doped Diamond (BDD) with intrinsic polycrystalline Diamond as passivation [7], which resulted in the first ever full diamond implant.
Micro-implants as shown in Figure 5, are usually made of three different parts. It consists of a substrate, conducting part and a non-conducting part, often referred to as passivation layer.
Figure 4 : 2-D cross-sectional view of the Full Diamond implant proposed by NEURODIAM [8]
Figure 5 : Simplified 2-D cross sectional view of an implant [8]
5

## Slide 8: Focus of this thesis
Research Objectives
To analyze the nature and function of conducting materials in this case specifically gold, along with various non-conducting parts also called passivation layers.
Distinguish behavior differences among IDE configurations based on impedance data and equivalent circuit modeling.
To conduct multiple electrical and electrochemical tests to gather data for a comprehensive understanding of these device’s performances.
Assess durability and stability enhancements provided by different passivation layers under various environmental conditions.

## Slide 9: Conducting and Non-Conducting Materials
Key Materials and Techniques
Conducting
Materials
Key material characteristics we look for is bio-compatibility, and stability.
Generally, Inert materials like platinum, gold, palladium, certain Silicon-based materials, metallic derivatives, Non-conventional conducting materials that were not initially developed for neural implants are now favored. [3][9]
In my thesis, Gold (Au) was chosen as the electrode material due to its exceptional biocompatibility, high conductivity, and long-term stability, despite the growing interest in alternative materials.
Why Gold ?
When inserted inside the body, it does not create adverse or toxic reactions upon encountering the skin, tissue, or internal organs and hence reduces postoperative risks.
Gold implants also show malleability, which is perfect for fabricating supple/flexible implants, high electric conductivity and reliability.[9]
Among all biocompatible metals, Gold shows the best corrosion resistance, and is high in density [10].

## Slide 10: Conducting and Non-Conducting Materials
Key Materials and Techniques
Non-conducting
Materials
What is a Passivation Layer?
A passivation layer is usually a non-conductive coating layer on top of the structure by treating or coating a metal surface to reduce its chemical reactivity so that it is less readily affected or corroded by the environment..[3][11], also in order to greatly improve its electrical characteristics.[12]
For all implantable devices, the biocompatibility of packaging and substrate materials is a prerequisite that must be met not only for the device’s long-term stability but also for the user’s safety.
In this thesis, two types of passivation materials, namely organic and inorganic passivation, were used over the gold electrode.
Organic Passivation- SU-8.
Inorganic Passivation- Silicon Nitride and Silicon Dioxide.

## Slide 11: Interdigitated Electrode - IDE
Key Materials and Techniques
1
2
3
The IDEs based sensor consists of two electrodes in which contains number of individual fingers like digits and are arranged in a comb-like structure to form micro-size gap between each electrode[9]. They can be fabricated using the well-known technique called photolithography.
This specific type of electrode was designed as a means to understand and improve the quality of passivation.
Applying a voltage to IDEs, either AC or DC, creates an electric field
between the fingers. The electric field lines generated by the electrodes penetrate the MUI and will change the impedance of the whole structure.
Figure 6: IDE with a sensing membrane as a parallel plate capacitor [14]
Figure 7: Coplanar interdigital sensor [14]
Figure 6 : The Conventional Interdigitated electrode (IDE) [9]
+V
-V

## Slide 12: Electrochemical Impedance Spectroscopy-EIS
Key Materials and Techniques
1
2
3
4
Electrochemical impedance spectroscopy (EIS) is a powerful technique that provides information on the processes occurring at the electrode-electrolyte interface[15].
A small signal excitation is applied to measure the impedance response. The electrochemical cell exhibits a pseudo-linear response, where a phase shift occurs, and the current response to a sinusoidal potential is also a sinusoid at the applied frequency. . Thus, the excitation signal is presented as a function of time (Eqn1).
EIS has the ability to study intrinsic material properties or specific processes that could influence conductance, resistance, or capacitance of an electrochemical system.
For this thesis, EIS was implemented to analyze the behavior of an IDE using an impedance analyzer with a small excitation voltage. The impedance results could be interpreted using Bode plots or Nyquist plots.
…..(1)
Figure 8 :Simplified EIS diagram[16]
Figure 9: Phase shift in current signal with reference to the applied voltage.[16]

## Slide 13: Equivalent Circuit Fit
Key Materials and Techniques
1
2
One of the key advantages of EIS in investigating the electrical and electrochemical characteristics is the direct association of the behaviour of real system to that of a simplified model., That is impedance data is fit to an equivalent circuit.
One of the most common circuits employed circuit for interpreting EIS experimental outcomes is the Randle’s Model that is illustrated in figure . Consisting of Rs,Cdl,Rct and W.
, where 1 means ideal capacitor.
Figure 10: Randles equivalent circuit model [17]
Figure 11: Bode Plot of the impedance magnitude of the Randle’s equivalent model [18]
Figure 12: Electro chemical reactions at the electrode/electrolyte surface[18]
……(2)

## Slide 14: Methodology
Data Analysis Approach
01
03
04
02
Phase 02 A. COMSOL Simulation
B. Experimental Analysis
C. Comparative Analysis
Phase 01 Microfabrication
Phase 03 Equivalent Circuit Fitting
Phase 04 Longevity Test

## Slide 15
MICROFABRICATION PROCESS

## Slide 16: Mask
Microfabrication Process
1
2
3
The stencil used to repeatedly generate a desired pattern on resist-coated substrates is called a mask[19]. The same mask used in the previous work was used which was designed using software CoventorWare10.4.
This mask was used to fabricate the Interdigitated electrodes with Gold as the conductive layer and SU8/Silicon dioxide/Silicon nitride as the passivation layer.
IDE with different fingers were fabricated. I.e., IDE-25, IDE-50 and IDE-100.
Figure 13 : (a) Top view depicting the dimensions of the IDE with respect to the mask used. (b) 2D- cross sectional view of the IDE without passivation layer.
(a)
(b)

## Slide 17: Mask
Microfabrication Process
Figure 14 : (a)The mask designed for fabrication of IDE with different pairs of fingers as viewed in Klayout.(b)Zoomed-in vision of a single strip with 25 pairs of fingers.(c) zoomed-in view of the finger-like structures.
(a)
(b)
(c)
(a)
(b)
(c)
(d)
Figure 15: (a),(c) 2D representation of the top view of the IDE with passivation layer.(b),(d) 2D cross-sectional view of an IDE with passivation layer. In (a)&(b) the passivation is deposited only on top of the electrodes leaving spaces to avoid the effect of water retention.

## Slide 18: Cleanroom Process
Microfabrication Process
Spin coat- SU-8
PECVD- Silicon Nitride/Silicon Dioxide

## Slide 19: Cleanroom Process
Microfabrication Process
Figure 16 : Top and Side view of the fabricated IDE
(a)
(b)
Passivation Layer (Transparent Part)
Metal (Gold)
Figure 17 : Microscopic view of gold covered with passivation

## Slide 20
EXPERIMENTAL ANALYSIS

## Slide 21: Experimental Setup and Test Conditions
Experimental Analysis
1
2
3
For the experimental Analysis, the impedance analyzer HP-4194A was used that is shown in Figure 20 , which analyzes impedance by scanning a specified frequency range, here, from 100Hz to 20MHz.
The fabricated IDE sample was connected to the impedance analyzer via PCB or wires, as shown in the Figure 18. The sample was first analyzed in air and then analyzed again after being immersed in a container with aqueous solutions as shown in Figure 19 .
Once calibrated, the HP-4194A automatically measures the impedance at each frequency point and plot the Bode plots on the screen. Once the sweep is complete, one can save the measurement data by data transfer to a connected PC using an appropriate interface.
Sample connected to the PCB.
Sample that was fabricated in the cleanroom.
Sample connected to the wires and its contacts encapsulated by copper and Kapton tape.
Figure 18: Fabricated IDE connected via PCB or wires for Impedance Analysis.
(a)
(b)

## Slide 22: Experimental Setup and Test Conditions
Experimental Analysis
Sample is connected via PCB or wire
Magnitude & Phase
Figure 20: Impedance Analyzer 4194A, (Input =1V), Freq range=(100-40M)Hz
Figure 19: Sample connected to the PCB immersed in aqueous solution.
Sample connected to the PCB.
Sample that was fabricated in the cleanroom.
Sample connected to the wires and its contacts encapsulated by copper and Kapton tape.
Figure 18: Fabricated IDE connected via PCB or wires for Impedance Analysis.
(a)
(b)

## Slide 23
COMSOL SIMULATIONS

## Slide 24: Simulation Setup and Parameters
COMSOL Simulation
1
2
There exists a vertical plane of symmetry located between the working and counter electrode. The interdigitated electrode digits of width W were separated by a distance S. The simulation domain for a working electrode in the presence of a mirror-image counter electrode is shown in Figure.
A frequency range of 100 Hz to 15 MHz is used in COMSOL for simulation. Initial simulations were conducted with air as the medium, followed by simulations in aqueous solutions, with and without passivation layer.
Figure 22 : Model simulated on COMSOL
Figure 21: Coplanar interdigital sensor [14], [Highlighted area]- to be modelled on COMSOL

## Slide 25: Simulation Setup and Parameters
COMSOL Simulation
Figure 24: Electric field lines penetrating the MUT
Figure 23 : Simulation modelled for 2 electrodes which is then divided by (N-1) number of fingers to obtain plots for N fingers.
σ , εr

## Slide 26: Simulation Setup and Parameters
COMSOL Simulation
Table 1: COMSOL parameters
Knee point/Relaxation frequency :
Table 2: Relaxation frequency of the media
……..(3)

## Slide 27: Simulation Setup and Parameters
COMSOL Simulation
Figure 25: (a)(b) Model of IDE to be simulated using COMSOL with varying thickness of passivation layer. (c)First order equivalent circuit of FEM model without passivation layer
(a)
(b)

## Slide 28
EXPERIMENT & SIMULATION
RESULTS

## Slide 29: No Passivation – In Air
COMSOL Simulation Results
Experimental Results

## Slide 30: No Passivation – In Air
Equivalent Curve Fit
Figure 26: (a)Equivalent circuit of IDE without passivation layer in air.(b) Simplified circuit in air.
(a)
(b)
Evaluated using :
Direct Measurement
Mask referral
~20ꭥ

## Slide 31: No Passivation – In all media
COMSOL Simulation Results
Experimental Results

## Slide 32: No Passivation – In all media
Equivalent Curve Fit
Figure 27: (a)Equivalent circuit of IDE without passivation layer in aqueous solutions.(b) Simplified circuit in solution.
High-frequency Path
Low-frequency Path
(a)
(b)
F < 100 Hz
ZCg=1/ωC

## Slide 33: No Passivation
Equivalent Curve Fit Results
High-frequency Path
Low-frequency Path
Figure 28: Curve fitting using Zview
Table 3: Electrical parameters of each medium
Table 4: Values of circuit elements in various media

## Slide 34: With passivation layer – In all media
COMSOL Simulation Results
Experimental Results

## Slide 35: With passivation layer – In all media
Figure 29: (a)Equivalent circuit of IDE with passivation layer in aqueous solutions.(b) Simplified circuit in solution.
(a)
11.24 kHz
65.16 kHz
19.98 MHz
11.24 kHz
65.16 kHz
19.98 MHz
Equivalent Curve Fit
(b)

## Slide 36: IDE With a Passivation Layer
Equivalent Curve Fit Results
Table 5: Relaxation frequency of the media
Table 6: Values of circuit elements in various media
Figure 29: Curve fitting using Zview

## Slide 37: Theoretical Capacitance of IDE
1
2
3
4
Figure 30 shows a layout of the electrode plane and a schematic diagram of the cross-section of an interdigital capacitor sensor with two interpenetrating comb electrodes. The gaps between electrodes have a width of G while the fingers have a width of W.
Each comb electrode is connected to a fixed potential (either +V or −V) and has a number of fingers of length L. The electrodes, made of gold in general, are deposited on a substrate whose thickness is much greater than the IDC-S spatial wavelength (λ = 2(W + G))
One can neglect the thickness of the electrodes (∼50 nm) since they are much thinner than their width (>50m),The metallization ratio is thus
Figure 31: The equivalent circuit for evaluation of the static capacitance of a semi-infinite top layer in a periodic IDC with 6 electrodes[`21].
Figure 30 : Layout of electrode plane [19].
The capacitance of one single semi-infinite layer can be evaluated as a function of two types of capacitances- (1) CI—being half the capacitance of one interior electrode relative to the ground potential and (2) CE—the capacitance of one outer electrode relative to the ground plane next to it. Using network analysis, the total capacitance is
………(4)
……..………(5)

## Slide 38: Theoretical Capacitance of IDE
Using conformal mapping technique [19], a typical IDE with a thick substrate of relative permittivity εS and one sensitive layer of relative permittivity ε1, will have their total capacitances CI,IDE and CE,IDE given by
……………….. (6)
……………….. (7)
Where, CI,air is capacitance of Air. CI,1 is the capacitance due to the sensitive layer and CI,S is the capacitance of the substrate. K(k) and K(k’) are elliptical integrals of the first kind , solved using MATLAB.
Table 7: Capacitance of IDE with varying fingers without passivation layer
Table 8: Capacitance of IDE-100 with / without passivation layer in air.

## Slide 39
COMPARATIVE ANALYSIS

## Slide 40: COMSOL Simulation V/S Experimental Analysis
Comparative Analysis
Due to the surface roughness and impurities, we obtain a CPE for the measured sensor, compared to a real capacitance due to the C in FEM simulations.[20]
1
2
Plot shows a comparative analysis of IDE with passivation in air. The difference in the slope of the impedance is that the low-frequency region can be explained due to the imperfect electrode surface of the real sensor compared to the simulation model. [20].

## Slide 41: IDE with Varied Finger Configurations
Comparative Analysis
Impedance increases with decrease in number of fingers [20]

## Slide 42: With Passivation V/S Without Passivation layer
Comparative Analysis

## Slide 43: Si3N4(1μm) V/S Si3N4(0.5μm)
Comparative Analysis

## Slide 44
LONGEVITY TEST

## Slide 45: Test Setup and Conditions
Longevity Test
1
2
3
4
To test the durability and performance of the passivation layer over time , samples were first enclosed within a resin cap produced by 3D printing and then submerged in 0.01M of PBS while set on heat. Increasing the heat over 37°C accelerated the speed-up aging of the samples.
To achieve reliable measurement with the long-life setup, an important step was making a hermetic setup where the electrodes would be kept at a constant temperature to increase the effect of heating for a long period.
A very well-established approach is to assume that the rate of aging is increased by a factor of 2ΔT/10, where ΔT is the temperature increase. This result is a mathematical expression of the empirical observation that increasing the temperature by about 10°C roughly doubles the rate of aging of the substance used .So, the experimental set-up was kept at a constant heat of 80°C.
To check the integrity of the sample we measure impedance regularly from 100 Hz to 10 MHz. The data is collected at 1 hour intervals for a period of over 30 days. A monitor is connected to this system to gather the data from the Raspberry Pi for analysis.
Sample enclosed within a resin cap
PDMS for better hermeticity
Thermocouple
Figure 32: Samples prepared for the longevity test in a cleanroom environment.

## Slide 46: Initial Measurements before Aging
Longevity Test

## Slide 47: Test Setup and Conditions
Longevity Test
Multimeter connected to the thermocouple that is placed inside the jar to measure the temperature inside the jar with samples.
Four Samples submerged in a jar filled with PBS 0.01M. The glass jar was then placed inside a hollowed-out metal block and the block was filled with biological oil to ensure heat transfer all around.
Hotplate where the temperature is set to about 80°C
Analog Discovery2 connected to a homemade multiplexer that has 4 ports. One port connected to one sample.
User Interface with Rasberry Pi connected to a monitor
Figure 33: System placed inside a Styrofoam box that is fully enclosed

## Slide 48: Measurements at T = 80°C
Longevity Test

## Slide 49: Measurements at T = 80°C
Longevity Test

## Slide 50: Post-Aging measurements in PBS 0.01M
Longevity Test

## Slide 51: Post-Aging measurements in PBS 0.01M
Longevity Test

## Slide 52: Post-Aging Calculations
Longevity Test
Table 11: Capacitance and resistance values of IDE in PBS solution, before and after aging.
Table 10: Capacitance and resistance values of IDE in air, before and after aging.
Table 9 : Percentage change in impedance, post aging.
2.6 times
2.98 times
4.15 times
4.09 times

## Slide 53
CONCLUSIONS & FUTURE WORK

## Slide 54: Conclusion and Future Work
The study successfully fabricated Interdigitated Electrodes (IDEs) using photolithographic processes, both with and without passivation, in a cleanroom environment.
Experimental and simulated analyses were conducted in different mediums, starting with air and then moving to various aqueous solutions (DI water, tap water, low salt, high salt, PBS 0.01M). This extensive evaluation demonstrated the versatility and adaptability of IDEs in diverse environments, highlighting their potential for varied applications.
Equivalent circuits were designed, that enabled a detailed understanding of the behavior of different IDE configurations.
Longevity tests on all IDEs provided valuable insights into their performance over time, both before and after aging. The results indicated the durability and long-term functionality of the IDEs
For better comparison purposes, it would be beneficial to standardize the passivation thickness across all samples.
Previous work at ESIEE Paris has already established diamond as the optimal passivation material. With a robust COMSOL model developed in this thesis, future work could further optimize the model to analyze its effectiveness.
Conclusions
Future Works

## Slide 55: References
Hassler, C., Boretius, T. and Stieglitz, T. (2011), Polymers for neural implants. J. Polym. Sci. B Polym. Phys., 49: 18-33. https://doi.org/10.1002/polb.22169
Y. Lu, X. Liu, R. Hattori, C. Ren, X. Zhang, T. Komiyama, D. Kuzum, Adv. Funct. Mater. 2018, 28, 1800002. https://doi.org/10.1002/adfm.201800002
Yang, Weiyang & Gong, Yan & Li, Wen. (2021). A Review: Electrode and Packaging Materials for Neurophysiology Recording Implants. Frontiers in Bioengineering and Biotechnology. 8. 10.3389/fbioe.2020.622923.
J. Lee et al., “Neural recording and stimulation using wireless networks of microimplants,” Nat Electron, vol. 4, no. 8, Art. no. 8, Aug. 2021, doi: 10.1038/s41928-021-00631-8. .
“NEURODIAM – High density full diamond cortical implant for long life time implantation.” Accessed: Sep. 14, 2022. [Online]. Available: https://www.neurodiam.eu/
Tabassum Sarah, Umme. "Comparison of Electrodes with Conventional, and Non Conventional Biomaterials to Novel Full Diamond Electrodes for Biomedical Applications." PhD diss., Université Gustave Eiffel, 2023. .
1
5
2
6
3
7
4
8
“Neuron Anatomy”,Accessed:October17,2023[Online].Available : https://nl.vecteezy.com/vector-kunst/358962-diagram-van-neuron-anatomy
H. An, K. Bai, Y. Yi, H. An, K. Bai, and Y. Yi, The Roadmap to Realize Memristive ThreeDimensional Neuromorphic Computing System. IntechOpen, 2018. doi: 10.5772/intechopen.78986.

## Slide 56: References
"Passivation(chemistry)".Accessed:October17,2023[Online].Available : https://en.wikipedia.org/wiki/Passivation_(chemistry) , https://www.electro-glo.com/passivation-vs-electropolishing-what-are-the-differences/
"Passivation". Accessed: October 17,2023[Online].Available : https://kayakuam.com/products/passivation/
“SIE Data Fitting – How to Get Good Initial Values of Equivalent Circuit Elements”.Accessed: October 17,2023[Online].Available: https://www.metrohm.com/fr_fr/applications/application-notes/autolab-applikationen-anautolab/an-eis-007.html
“Electrochemcal Impedance Spectroscopy (EIS) Basics”.Accessed:October20,2023[Online.]Available:https://pineresearch.com/shop/kb/theory/eis-theory/eis-basics/
Márton G, Tóth EZ, Wittner L, et al. The neural tissue around SU-8 implants: A quantitative in vivo biocompatibility study. Mater Sci Eng C Mater Biol Appl. 2020;112:110870. doi:10.1016/j.msec.2020.110870
9
13
10
14
11
15
12
16
Khan, Md Rajibur Rahaman & Kang, Shin-Won. (2015). Highly Sensitive Multi-Channel IDC Sensor Array for Low Concentration Taste Detection. Sensors (Basel, Switzerland). 15. 13201-13221. 10.3390/s150613201.
E. S. Vikulova et al., “Application of Biocompatible Noble Metal Film Materials to Medical Implants: TiNi Surface Modification,” Coatings, vol. 13, no. 2, Art. no. 2, Feb. 2023, doi: 10.3390/coatings13020222.
Jun, Li et al. “Simulation of Interdigitated Electrodes (IDEs) Geometry Using COMSOL Multiphysics.” 2018 International Conference on Intelligent and Advanced System (ICIAS) (2018): 1-6.

## Slide 57: References
17
Rui Igreja, C.J. Dias,Analytical evaluation of the interdigital electrodes capacitance for a multi-layered structure,Sensors and Actuators A: Physical, Volume 112, Issues 2–3,2004,Pages291-301,ISSN0924-4247,https://doi.org/10.1016/j.sna.2004.01.040.
Mitra, S.K., & Chakraborty, S. (Eds.). (2012). Microfluidics and Nanofluidics Handbook: Fabrication, Implementation, andApplications(1st ed.).CRC Press https://doi.org/10.1201/b11188
Korek, Eva Maria ; Teotia, Reva ; Herbig, David et al. / Electrochemical Impedance Spectroscopy for Ion Sensors with Interdigitated Electrodes : Capacitance Calculations, Equivalent Circuit Models and Design Optimizations. In: Biosensors. 2024 ; Vol. 14, No. 5.
18
19
21
20
M. S. A. Rahman, S. C. Mukhopadhyay, and P.-L. Yu. «Novel Planar Interdigital Sensors». In: Novel Sensors for Food Inspection: Modelling, Fabrication and Experimentation. Springer, 2014, pp. 11–35. doi: 10.1007/978-3-319- 04274-9_2 (cit. on pp. 52, 53).
Hend S. Magar, Rabeay Y. A. Hassan, and Ashok Mulchandani. «Electrochemical Impedance Spectroscopy (EIS): Principles, Construction, and Biosensing Applications». In: Sensors 21.19 (2021). doi: 10 . 3390 / s21196578

## Slide 58
THANK YOU !
Pr. Gaëlle LISSORGUES
(Dr,HDR)
Industrial Supervisor
Pr. Lionel ROUSSEAU
Pr. Matteo COCUZZA
(Dr,HDR)
Industrial Co-Supervisor
Academic Supervisor